## Feuil1
Sample preparation
Sample preparation | 1 inch epoxy plug
Imaging | Optical
LASER ablation system U-Pb dating
Make, Model & Type | ESI, NWR193
Ablation cell & size | TV2 150mmx150mm
LASER wavelength (nm) | 193
Pulse duration (ns) | 4
Fluence (J.cm-2) | 0.9 - 1.4
Repetition rate (Hz) | 15
Spot size (µm) | 150
Sampling | Spot
Carrier gas | He 5.0 ~0.9 l.min-1, Ar make-up gas ~1 l.min-1 combined using a Y-piece ~50cm before the connection to the injector
Ablation duration (s) | 3s warmup - 4 s preablation - 15 s washout - 3s warmup - 20 s ablation - 15 s washout
ICP-MS system U-Pb dating
Make, Model & Type | Thermo Fisher Scientific, Element XR
RF Power (W) | 1200 - 1300
Make-up gas flow (l.min-1) | ~1
Detection System | SEM forced in counting mode
Masses measured | 206Pb, 207Pb, 208Pb, 232Th, 238U
Integration time per peak (s) | 0.04, 0.04, 0.02, 0.012, 0.012
Total integration time per cycle (s) | 0.129 // Mass accuracy option
Sensitivity | ~100 kcps for 823 ng/g U
IC dead time (s) | 26
Data processing U-Pb dating
Gas blank | on-peak zero substraction, ~7s signal intensity integration during the washout between each preablation and ablation
Calibration | Daily tuning with NIST-612 then during sequence analyses use of NIST-614 glass standard as primary reference material for drift (Woodhead and Hergt, 2001); WC-1 carbonate reference material for matrix matching of 206Pb/238U; AUG-B6 as secondary material; after 2022 use of JCP1-np from myStandards GmbH for [U], [Pb], [Th] calibration
Reference material info | WC1 (Roberts et al. 2017) and AUG-B6 (Pagel et al. 2018)
Data processing | Raw intensities, baseline subtraction, calulation of log-ratios with Iolite 3.65 (Paton et al, 2011) Instrumental drift check, geometric averages of the ratios and 206Pb/238U ratios corrections using a fractionation factor calculated from WC1 radiogenic intercept with the concordia, exclusion of spots with mean 207Pb intensities using a cutoff value calculated by adding ten times the standard deviation to the average baseline intensity on 207Pb for each session using an in-house Excel spreadsheet IsoplotR (Vermeesch, 2018) for T-W isochrons (model-1), 2e as input, 2se as output, intercept ages and initial Pb compositions
Mass discrimination | 1 set of standards every ~25 spots
Common-Pb correction, composition and uncertainty | None applied. Ages calculated from regressions used in Tera Wasserburg plots
Uncertainty level & propagation | Ages are quoted at 2se with overdispertion including the propagation of systematic uncertainty by quadratic addition of 2.5% on WC1 age and an extra 3.5% to account for the uncertainty on AUG-B6 secondary standard which is still under evaluation through an inter laboratory comparison
Quality control / Validation | AUG-B6 - mean age: 43 Ma±2 Ma (2SD, n=3) in agreement with the age published by Pagel et al (2018): 43Ma±1.5Ma